Risk Assessment Form

Cardiff School of Biosciences

IMPORTANT: Before carrying out the assessment, please read the Guidance Notes

1. General Information

| Department | Cardiff School of Biosciences | Building | Sir Martin Evans (BIOSI 2) | Room No | East 3.15 |
| --- | --- | --- | --- | --- | --- |
| Name of Assessor | Anthony J Hayes | Date of Original Assessment | 18/01/2017 | Assessment No or practical module No | 5 |

Status of Assessor:  Staff 	Postgraduate 	Undergraduate 	Other:

(Specify)

2. Brief Description of Procedure/Activity including its Location and Duration

Operation of spinning disc confocal microscopes, employing class 3B lasers, within live cell imaging suite, Cardiff School of Biosciences (BIOSI 2; E/3.15). Spoke system administrator: Dr Peter Watson (ext 79042). This risk assessment covers use of the spinning disc confocal system housed in a live cell imaging suite in BIOSI 2; E/3.15, system administrator: Dr Peter Watson. The system consists of an Olympus IX71 motorised inverted microscope with a Crest Optics X-Light V2 confocal imager and a Cairn Research tri-line laser bank (Cairn ref# 6737), power unit (Cairn ref# 6738) and control rack (Cairn ref# 6739). This risk assessment should be read in conjunction with the risk assessment for fluorescence microscopy. The system belongs to laser hazard class 3B and is designed, manufactured, and maintained to comply with applicable performance standards for Class 3B laser products. Laser specifications: Laser modules; Power; Lines. OBIS Diode 405 cw; 100mW; 405nm. OBIS Diode 488 cw; 150mW; 488nm. OBIS Diode 561 cw; 100mw; 561nm. cw, continuous wave The lasers are connected via a shielded fibre to the confocal imaging head. Lasers are activated via a key switch on the control module and via software. Beam paths are shielded except for a fixed open beam between objective and specimen. Warning labels have been affixed near all apertures or removable parts where exposure to laser light is possible. Warning notices and safety guidance notes have been positioned around the microscope for maximum visibility to the user.

2a. Is your work governed by specific legislation i.e:

(Tick as appropriate, see guidance notes)

Human Tissue (HTA-work involving human tissue):				Approval compliance obtained

GM (any genetically modified organism including plant and animals):		Approval compliance obtained

Radiation (radioisotopes, sealed sources):					Approval compliance obtained

Controlled Drugs:								Approval compliance obtained

Non ionising radiation (lasers, magnetism):					Approval compliance obtained

Use of human subjects (Ethics):							Approval compliance obtained

3. Persons at Risk	Are they	Notes

| Staff Visitor Contractor Students Vulnerable groups |  | Trained Disabled Inexperienced Competent |  | Trained staff/students (unsupervised) and staff/students undergoing training (supervised) are potentially at risk. Untrained users are not permitted access to the instrument. |
| --- | --- | --- | --- | --- |

4. Level of Supervision			Notes

| None Constant Periodic Training Required | Training is required for independent usage of the confocal microscope. Untrained users are not permitted independent access to the instrument. |
| --- | --- |

5. Will Protective Equipment Be Used?  Please give specific details of PPE

| Head Eye Ear Body Hand Foot | For live cell imaging applications (see below), the use of a lab coat and protective gloves are essential. Safety goggles would severely hamper routine use of the microscope but may be necessary for some live cell imaging applications (see below). |
| --- | --- |

6. Is the Environment at Risk?		Notes

| Yes No | Specimens should be sealed before observation under the microscope and handled with care using the appropriate level of protection (see above). Leaky or cracked samples must not be examined. Samples on glass histology slides should be mounted under coverslips with hard-set mountant or sealed with nail varnish to prevent leakage. Live cells grown within tissue culture plastics should be sealed with Parafilm to prevent spillage. Any spillages should be cleaned up immediately and the area swabbed with 95% alcohol. Broken glass slides should be disposed of in the contaminated sharps bin – broken fragments of glass should be brushed on to paper and disposed of in a similar fashion. |
| --- | --- |

7. Will Waste be generated?		If ‘yes’ please give details of disposal

| Yes No | All biological samples and hazardous waste must be appropriately disposed of by the user. Potential risks associated with the waste material should be addressed by the user in a separate risk assessment form. |
| --- | --- |

8. Hazards involved

| Work Activity / Item of Equipment / Procedure / Physical Location | Hazard | Control Measures and Consequence of Failure | Likelihood (0 to 5)  | Severity (0 to 5) ═ | Level of Risk |
| --- | --- | --- | --- | --- | --- |
| Laser scanning confocal microscopy | Class 3B laser radiation; Eye strike | 1. Physical measures: The spinning disc confocal microscope is housed within a low light microscopy suite (BIOSI 2; E/3.15) with clearly labelled signage to prevent unauthorised access. Laser power is adjusted to the minimum necessary to produce a signal. Laser light is fed through an optical waveguide and therefore shielded until it leaves the microscope objective and reaches the specimen. The illumination mode controller on the front of microscope stand switches between confocal and eyepiece observation. 2. Administrative measures: Safety notices and precautions on and around the microscope should always be observed before and during operation. Standard safety procedures are compulsory for all users and must be incorporated into all experimental protocols. In summary, users should: Never look directly into the laser beam or a reflection of the beam, Never interfere with the beam path, Never change a specimen during scanning, Never change objectives during scanning, Never change filter cubes during scanning. A register of users is maintained, and the use of the equipment is restricted to: Trained staff (unsupervised) Staff under training (supervised). Training includes safety provisions based on information supplied by Lasermet, OSHEU and AURPO. A laminated emergency information card in E/3.15 provides information for emergency medical treatment. | 1 | 4 | 4 |
|  | Broken glass | Broken glass from histology slides or coverslips should be disposed of in the contaminated sharps bin. | 1 | 2 | 2 |
| Observation of live cells/tissues | Biological contaminants | For individual research application involving live cells/tissues, appropriate control measures should be in place to reflect the potential risk of the organism/tissue under study. These must be evaluated by the user and the resultant risk assessment approved by the system administrator, Dr Pete Watson & Safety representative where necessary. A copy of the risk assessment should be filed within the facility and subject to periodic review. All samples should be clean and sealed to prevent leakage/spillage during observation on the confocal microscope. Leakages and spillages should be cleaned up immediately and the area swabbed with 95% alcohol. All live samples should be removed from room E/3.15 after microscopical observation and disposed of appropriately by the user. | * | * | * |
| Fluorescence microscopy (Refer to additional Fluorescence microscopy risk assessment) | Mercury lamp burst | Administrative measures to monitor lamp usage (record of lamp usage) and replacement of Hg lamps approaching the end of their lifespan (i.e., before 300hrs). In the extremely unlikely event of a lamp burst, the room will be vacated immediately. | 1 | 3 | 3 |

* application dependent (requires additional risk assessment where necessary)

9. Chemical Safety (COSHH Assessment)

| Hazard | Control Measures | Likelihood (0 to 5)  | Severity (0 to 5) ═ | Level of Risk |
| --- | --- | --- | --- | --- |
| Seepage of mountant from broken or leaky sample preparations. Aqueous mountants often contain DNA-binding probes (e.g., DAPI; 4',6-diamidino-2-phenylindole) used for counterstaining of cell nuclei. | Users should use hard-set mountant, or seal sample preparations with nail varnish to prevent leakage. Gloves recommended for use with aqueous mountants. Sample preparations should be handled with care and disposed of appropriately (see above). | 1 | 2 | 2 |
| Lens cleaning solution containing 90% petroleum ether and 10% isopropanol (working volume of 250ml) | Users should take care to avoid any spillages when cleaning microscope objective lenses. Any spillages should be reported immediately to The Facility Manager. No more than a 250ml working volume should be made up at any one time in a sealed, clearly labelled container. All solvent stocks should be kept in a locked solvent cabinet. | 1 | 2 | 2 |
| The immersion oil used for microscope objectives is an irritant | Users should take care to avoid contact with the immersion oil. Wash contaminated skin with soap and water. | 1 | 1 | 1 |

Scoring Criteria for Likelihood (chance of the hazard causing a problem) Likelihood

Scoring Criteria for Severity of Injury (or illness) resulting from the hazard

Likelihood

|  | 5 | Almost Certain | 5 | 10 | 15 | 20 | 25 |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | 4 | Very Likely | 4 | 8 | 12 | 16 | 20 |
|  | 3 | Likely | 4 | 6 | 9 | 12 | 16 |
|  | 2 | Unlikely | 2 | 4 | 6 | 8 | 10 |
|  | 1 | Very Unlikely | 1 | 2 | 3 | 4 | 5 |
| Severity |  |  | No Injury / Illness | First Aid Required | Minor Injury | Major Injury | Death |
|  |  |  | 1 | 2 | 3 | 4 | 5 |

Score Action to be taken:

0-5	Low Risk	No further action needed.

6-9	Medium Risk	Appropriate additional control measures should be implemented

10-25	High Risk	Additional control measures MUST be implemented. Work MUST NOT commence until such measures are in place. If work has already started it must STOP until adequate control measures are in place

10. Source(s) of information used to complete the above e.g., Supervisor, Web etc….

Lasermet laser safety course notes; AURPO guidance Note 7: Guidance on the safe use of lasers in education and research; Cardiff University OSHEU Guidance document NIRP2: Working with lasers and other optical radiations. Laser safety training, Cardiff University, Online module, 2018.

11 Additional Control Measures – Likelihood and Severity are the values with the additional controls in place

| Work Activity / Item of Equipment / Procedure / Physical Location | Hazard and Existing Control Measures | Additional Controls needed to Reduce Risk | Likelihood (0 to 5)  | Severity (0 to 5) ═ | Level of Risk |
| --- | --- | --- | --- | --- | --- |

After the implementation of new control measures the procedure/activity should be re-assessed to ensure that the level of risk has been reduced as required.

12. Action in the Event of an Accident or Emergency

Report immediately to system administrator, Dr Peter Watson (ext 79042). Follow OSHEU guidance (ext 74910). Non-emergency medical attention call 111; medical emergency call 999. Cardiff eye unit (direct line): 02920 743862; Cardiff A&E (direct line): 02920 748025/8031. Provide details of the laser(s) in use to medical staff.

13. Arrangements for Monitoring the Effectiveness of Control

Ad-hoc visual checks and periodic review of existing risk assessments. School Safety inspections, internal and external safety audits.

14. Review:  This assessment will be reviewed following any changes to procedure or three years from last review date.

| Name of Reviewer: | Dr Anthony J. Hayes |  | Date of latest Review: | 29/06/2023 |
| --- | --- | --- | --- | --- |
| Have the Control measures been effective in controlling the risk? | Yes |  |  |  |
| Have there been any changes in the procedure or in information available which affect the estimated level of risk? | Relocation of microscope to new site: BIOSI 2; E/3.15; risk assessment updated accordingly (5.4.22) Updated contact details for accident or emergency (29/06/2023) |  |  |  |
| What changes to the Control Measures are required? | None |  |  |  |

15. Signatures for printed copies:

| Assessor: | Dr Anthony J. Hayes | Signature: | AJH | Date: | 18.1.17 |
| --- | --- | --- | --- | --- | --- |
| Reviewed by: | Dr Anthony J. Hayes | Signature | AJH | Date: | 29.1.2019 |
| Reviewed by: | Dr Anthony J. Hayes | Signature: | AJH | Date: | 14.2.2020 |
| Reviewed by: | Dr Anthony J. Hayes | Signature: | AJH | Date: | 5.4.2022 |
| Reviewed by: | Dr Anthony J. Hayes | Signature: | AJH | Date: | 29/06/2023 |
| Person involved on risk assessment or issued to | Risk assessment must be read and understood by all new microscope users during induction. A copy will be available online via SOP repository. Hard copies in lab and office. | Signature |  | Date: |  |